## Sheet1
APEX RF Monitor Cable Calibrations
Signal Name | Cable # | Chassis, RF Board # and Channel # | Signal Source Power Level (dBm) | Signal Source Power Level (Watts) | Signal Source Coupling Factor (dB) | Cable Atten (dB) | Splitter Atten (dB) | Cable Length (°) | Splitter & Cable Length (°) | Expected F.S. Signal Level @ Chassis R.P. connector (dBm) | Test Signal Level @ Chassis R.P. connector (dBm) | Chassis Signal Path Loss (back calc'd) (dB) | Chassis Signal Path Length Difference (°) | Chassis Atten (dB) | Total Atten & Loss (dB) | Measured Signal Level at ADC (dBFS) | Signal level at F.P. Port for given Test Signal (dBm) | Coupling Factor for Internal Coupler (dB) | Front Panel Port Coupling (from R.P.) (dB) | Front Panel Port Coupling (from source) (dB) | Overall Full Scale (Watts) | Real Full Scale for Input at -1.72 dB (Watts)
Cav A3 FWD, (splitter) | 18 | llrf1,1,2 | 77.782 | 60006.7354 | -50 | -3.727 | -3.312 | 47.09 | 11.1 | 20.743 | 20.59 | -15.46 |  | -6.05 | -78.549 | -0.92 | -5.9 | -20.44 | -26.49 | -83.529 | 71597.8531
Cav A3 REV | 17 | llrf1,1,3 | 77.782 | 60006.7354 | -50 | -3.8 | 0 |  |  | 23.982 | 24.21 | -15.725 |  | -8.935 | -78.46 | -0.45 | -5.2 | -20.475 | -29.41 | -83.21 | 70145.5298
Cav Cell Probe #1 | 16 | llrf1,1,4 | 80.792 | 120005.182 | -47.8 | -3.7 | 0 |  |  | 29.292 | 27.57 | -16.347 |  | -12.963 | -80.81 | -1.74 | -5.9 | -20.507 | -33.47 | -84.97 | 120503.594 | 81096.1058
Laser Freq Input |  | llrf1,1,1 | [8 empty] | -16.1 | -14.85 |  | 30 |  | -0.95 | -6.3 | -20.2 | 9.8
Cav A4 FWD | 21 | 2,1,1 | 77.782 | 60006.7354 | -50 | -3.706 | 0 | 131.75 |  | 24.076 | 24.13 | -2.183 |  | -22.477 | -78.366 | -0.53 | -18.8 | -20.453 | -42.93 | -96.636 | 68643.5918
Cav A4 REV | 20 | 2,1,2 | 77.782 | 60006.7354 | -50 | -3.8 | 0 |  |  | 23.982 | 24.05 | -1.97 | 18.59 | -22.49 | -78.26 | -0.41 | -18.9 | -20.46 | -42.95 | -96.75 | 66988.4609
Cav Cell Probe #2 | 19 | 2,1,3 | 80.792 | 120005.182 | -50.13 | -3.7 |  |  |  | 26.962 | 26.93 | -2.556 |  | -25.414 | -81.8 | -1.04 | -19 | -20.516 | -45.93 | -99.76 | 151356.1248
Cav A3 FWD, (splitter) |  | 2,1,4 | 77.782 | 60006.7354 | -50 | -3.727 | -3.308 | 47.09 | 11.2 | 20.747 | 20.78 | -1.703 |  | -20.097 | -78.835 | -1.02 | -19.7 | -20.383 | -40.48 | -97.515 | 76471.5688
Tetrode A3 FWD | 7 | 1,2,1 | 77.782 | 60006.7354 | -63.9 | -0.561 | 0 | -66.276 |  | 13.321 | 13.37 | -2.165 |  | -12.055 | -78.681 | -0.85 | -19.1 | -20.415 | -32.47 | -96.931 | 73807.4158
Tetrode A3 REV | 8 | 1,2,2 | 77.782 | 60006.7354 | -64 | -0.6 | 0 |  |  | 13.182 | 13.17 | -1.451 | -6.6 | -12.039 | -78.09 | -0.32 | -19.3 | -20.431 | -32.47 | -97.07 | 64416.9266
Tetrode A4 FWD | 9 | 1,2,3 | 77.782 | 60006.7354 | -63.6 | -0.611 | 0 | 134.84 |  | 13.571 | 13.55 | -1.9 |  | -12.11 | -78.221 | -0.46 | -19 | -20.44 | -32.55 | -96.761 | 66389.592
Tetrode A4 REV | 10 | 1,2,4 | 77.782 | 60006.7354 | -63.8 | -0.6 | 0 |  |  | 13.382 | 13.36 | -2.22 |  | -12.01 | -78.63 | -0.87 | -19.3 | -20.65 | -32.66 | -97.06 | 72945.751
Circ 1 Load FWD | 1 | 2,2,1 | 77.782 | 60006.7354 | -50 | -1.4 | 0 |  |  | 26.382 | 26.43 | -2.058 |  | -25.252 | -78.71 | -0.88 | -19.3 | -20.478 | -45.73 | -97.13 | 74301.9138
Circ 1 Load REV | 2 | 2,2,2 | 77.782 | 60006.7354 | -50.1 | -1.4 | 0 |  |  | 26.282 | 26.26 | -2.62 |  | -23.78 | -77.9 | -0.14 | -18 | -20.48 | -44.26 | -95.76 | 61659.5002
Circ 2 Load FWD | 5 | 2,2,3 | 77.782 | 60006.7354 | -50 | -1.4 | 0 |  |  | 26.382 | 26.36 | -2.88 |  | -23.58 | -77.86 | -0.1 | -17.7 | -20.48 | -44.06 | -95.46 | 61094.2025
Circ 2 Load REV | 4 | 2,2,4 | 77.782 | 60006.7354 | -50 | -1.3 | 0 |  |  | 26.482 | 26.43 | -2.294 |  | -24.456 | -78.05 | -0.32 | -18.5 | -20.474 | -44.93 | -96.23 | 63826.3486
SSPA A1 FWD | 15 | 1,1,1 | 64.772 | 3000.544 | -53.8 | -0.462 | 0 | -72.444 |  | 10.51 | 10.46 | -1.974 |  | -8.856 | -65.092 | -0.37 | -18.9 | -20.504 | -29.36 | -83.622 | 3229.9812
SSPA A1 REV | 14 | 1,1,2 | 64.772 | 3000.544 | -50 | -0.5 | 0 |  |  | 14.272 | 13.94 | -1.38 | -7.22 | -12.7 | -64.58 | -0.14 | -19.1 | -20.34 | -33.04 | -83.54 | 2870.7806
SSPA A2 FWD | 13 | 1,1,3 | 64.772 | 3000.544 | -53.7 | -0.524 | 0 | 60.345 |  | 10.548 | 10.25 | -1.675 |  | -8.705 | -64.604 | -0.13 | -19.1 | -20.645 | -29.35 | -83.574 | 2886.689
SSPA A2 REV | 12 | 1,1,4 | 64.772 | 3000.544 | -50 | -0.6 | 0 |  |  | 14.172 | 13.95 | -1.46 |  | -12.65 | -64.71 | -0.16 | -19.2 | -20.5 | -33.15 | -83.75 | 2958.0125
Splitter S -> 1, incl cable | [6 empty] | 3.312 | 11.1
Splitter S -> 2, incl cable | [6 empty] | 3.308 | 11.2
[13 empty] | Signal Name |  | Expected Signal Level at ADC (dB) | Measured Signal Level at ADC (dBFS) | ADC Channel Error (dB) | Expected Signal Level at ADC for Real Ful Scale Input (dBFS) | ADC Channel Error (dB)
Cav A3 FWD, (splitter) | [12 empty] | Cav A3 FWD, (splitter) |  | -0.767 | -0.92 | -0.153
Cav A3 REV | [12 empty] | Cav A3 REV |  | -0.678 | -0.45 | 0.228
Cav Cell Probe #1 | [12 empty] | Cav Cell Probe #1 |  | -0.018 | -1.74 | -1.722 | -0.02 | -0.002
Laser Freq Input | [12 empty] | Laser Freq Input
Cav A4 FWD | [12 empty] | Cav A4 FWD |  | -0.584 | -0.53 | 0.054
Cav A4 REV | [12 empty] | Cav A4 REV |  | -0.478 | -0.41 | 0.068
Cav Cell Probe #2 | [12 empty] | Cav Cell Probe #2 |  | -1.008 | -1.04 | -0.032
Cav A3 FWD, (splitter) | [12 empty] | Cav A3 FWD, (splitter) |  | -1.053 | -1.02 | 0.033
Tetrode A3 FWD | [12 empty] | Tetrode A3 FWD |  | -0.899 | -0.85 | 0.049
Tetrode A3 REV | [12 empty] | Tetrode A3 REV |  | -0.308 | -0.32 | -0.012
Tetrode A4 FWD | [12 empty] | Tetrode A4 FWD |  | -0.439 | -0.46 | -0.021
Tetrode A4 REV | [12 empty] | Tetrode A4 REV |  | -0.848 | -0.87 | -0.022
Circ 1 Load FWD | [12 empty] | Circ 1 Load FWD |  | -0.928 | -0.88 | 0.048
Circ 1 Load REV | [12 empty] | Circ 1 Load REV |  | -0.118 | -0.14 | -0.022
Circ 2 Load FWD | [12 empty] | Circ 2 Load FWD |  | -0.078 | -0.1 | -0.022
Circ 2 Load REV | [12 empty] | Circ 2 Load REV |  | -0.268 | -0.32 | -0.052
SSPA A1 FWD | [12 empty] | SSPA A1 FWD |  | -0.32 | -0.37 | -0.05
SSPA A1 REV | [12 empty] | SSPA A1 REV |  | 0.192 | -0.14 | -0.332
SSPA A2 FWD | [12 empty] | SSPA A2 FWD |  | 0.168 | -0.13 | -0.298
SSPA A2 REV | [12 empty] | SSPA A2 REV |  | 0.062 | -0.16 | -0.222